## Slide 1: Implant materials for bone and joints
BIEN 235
Dr. Teresa Murray
Louisiana Tech University
College of Engineering and Science

## Slide 2: Fixation and replacement joint devices
DePuy products advertisement
… putting Michaelangelo back together again …
Louisiana Tech University
College of Engineering and Science
http://www.youtube.com/watch?NR=1&v=IPRddQcfD9s

## Slide 3: Bone
Louisiana Tech University
College of Engineering and Science
Anisotropic Composite Biological Material
Compression: Strong
Tension: weaker
Torsion and bending:
Weak
This textbook figure has 100-500 mm and 3-7 mm. This should be microns (mm)

## Slide 4: Hip anatomy and arthroplasty
Illium
Acetabulum
Femoral head
Louisiana Tech University
College of Engineering and Science
http://www.millerandzois.com/stryker-hip-settlement.html

## Slide 5: Components of traditional hip implant
Hip implant types Animation Zirconium Advertisement
http://www.youtube.com/watch?NR=1&v=k-QOqOayBUQ
Louisiana Tech University
College of Engineering and Science

## Slide 6: Hip replacement surgery
Animation of total hip arthroplasty
Louisiana Tech University
College of Engineering and Science
http://www.youtube.com/watch?v=Bwi0i1EtbpI&feature=related

## Slide 7: Hip replacement components
Louisiana Tech University
College of Engineering and Science
What are the desired material properties?
Stem (1-2 parts plus femoral head)
Femoral head (may be 1 piece with stem)
Acetabular cup (may have >1 part)
What are the biological concerns?

## Slide 8: Components of traditional hip implant
Hip implant types Animation Zirconium Advertisement
Louisiana Tech University
College of Engineering and Science
http://www.youtube.com/watch?NR=1&v=k-QOqOayBUQ

## Slide 9: Some of the materials
Co Cr alloys and Ti alloys
Ceramics
Polymers (UHMWPE)
Bone cement
What are the mechanical properties?
How biocompatible are they?
Louisiana Tech University
College of Engineering and Science
http://www.youtube.com/watch?NR=1&v=k-QOqOayBUQ

## Slide 10: Stress shielding
Louisiana Tech University
College of Engineering and Science
http://www.bananarepublican.info/Stress_shielding.htm
NOTE: Each leg has half the stress from the whole body

## Slide 11
Date of download: 10/18/2011
Copyright © The Journal of Bone & Joint Surgery. All rights reserved.
From: Long-Term Clinical Consequences of Stress-Shielding after Total Hip Arthroplasty without Cement*
The Journal of Bone and Joint Surgery (American) 1997; 79:1007-12
Radiographs showing pronounced bone remodeling with evidence of bone ingrowth (square). Immediate postoperative radiograph (left) shows an anatomic medullary locking stem that fills the intramedullary canal at the isthmus. Radiograph (right) after 2 years is typical of patients in the study. Note pronounced bone-remodeling, marked thinning and decreased density of the cortices, and smooth surface of substrate that is indicative of osseointegration.
“Stress-shielding” eventually weakens bone on one side of implant
What happens then?

## Slide 12: Osteolysis
Metal-on-polyethylene implants
  microscopic polyethylene wear fragments shed from implant
  immune system attacks these
  can ‘eat away’ surrounding bone
  weakens bone, implants loosen
Many orthopedic surgeons blame osteolysis as the major cause of hip implant failure
http://orthopedics.about.com/od/hipkneereplacement/a/implants.htm
Louisiana Tech University
College of Engineering and Science
http://www.youtube.com/watch?NR=1&v=k-QOqOayBUQ

## Slide 13: Osteolysis
Newer ceramic-on-polyethylene, ceramic-on-ceramic, and metal-on-metal implants developed
Hope is they will reduce wear fragments and osteolysis
http://orthopedics.about.com/od/hipkneereplacement/a/implants.htm
Louisiana Tech University
College of Engineering and Science
http://www.youtube.com/watch?NR=1&v=k-QOqOayBUQ

## Slide 14: But are these always met?
Louisiana Tech University
College of Engineering and Science
Materials used in total hip replacement have 4 common characteristics
Biocompatible. Do not create a local or a systemic rejection response.
Resistant to corrosion, degradation, and wear. Retain strength and shape for years. Resist wear and prevent further destruction of bone caused by particulate debris (generated when implant parts move against each other).
Mechanical properties duplicate the structures they replace. Parts are (1) Strong enough to withstand weight-bearing loads, (2) flexible enough to undergo stress without breaking, and (3) able to move smoothly against each other.
Meet the highest standards. Fabrication and quality control at a reasonable cost
http://orthoinfo.aaos.org/topic.cfm?topic=A00355

## Slide 15: New material for implant
Hip implant types Animation Zirconium Advertisement
http://www.youtube.com/watch?NR=1&v=k-QOqOayBUQ
Louisiana Tech University
College of Engineering and Science

## Slide 16: New product development
Hip implant new materials being developed, such as incorporating carbon nanotubes.
http://www.azom.com/materials-video-details.aspx?VidID=264
Louisiana Tech University
College of Engineering and Science

## Slide 17: Growing bone: novel approaches
Louisiana Tech University
College of Engineering and Science
http://www.azom.com/materials-video-details.aspx?VidID=103

## Slide 18
Louisiana Tech University
College of Engineering and Science
http://abclocal.go.com/kabc/video?id=8312311&syndicate=syndicate&section=
Growing bone: novel approaches

## Slide 19: Spinal stenosis
Louisiana Tech University
College of Engineering and Science

## Slide 20: Spinal fusion
Louisiana Tech University
College of Engineering and Science

## Slide 21: Optional slides on ethical issues
You do not need to know this for an exam! However, this may be helpful for learning about technical and ethical issues regarding orthopedic implants.

## Slide 22: Off-label testing of bone cement
A manufacturer of bone cement went straight to clinical trials (no IDE), despite an indication that their bone cement caused blood clotting. Three deaths occurred and they're paying $20+ million in damages.
http://www.fda.gov/MedicalDevices/Safety/RecallsCorrectionsRemovals/ListofRecalls/ucm238423.htm
http://www.outpatientsurgery.net/news/2010/12/7-synthes-admits-guilt-in-illegal-off-label-bone-cement-testing
Louisiana Tech University
College of Engineering and Science

## Slide 23: DePuy MOM total hip recall
Hip animation and implant animation with metal debris (metalosis)
http://www.youtube.com/watch?v=nZXj7esOK7A&feature=related
Louisiana Tech University
College of Engineering and Science

## Slide 24
Louisiana Tech University
College of Engineering and Science
“Information for Physicians Who Believe They May Have a Patient With a Sensitivity to Metal
The most common metallic sensitizers that are used in orthopaedic alloys are nickel, cobalt, and chromium.
Orthopaedic grades of stainless steel and cobalt-chromium alloys both contain these materials while titanium and titanium alloys do not. … The excellent biocompatibility of titanium and its alloys, such as Ti-6Al-4V alloy, is well documented. Persons with a history of allergies, including sensitivities to cobalt, chromium, or nickel, generally do not exhibit or develop sensitivity to titanium or other constituents of Ti-6Al-4V alloy.”
http://www.zimmer.com/web/enUS/pdf/Metal_Sensitivity_Information_Sheet_02-21-2011.pdf

## Slide 25: Zimmer Durom acetabular cup recall
In US, coating on Zimmer “Durom ” cup is thicker than those sold outside the US. No revision surgeries reported outside the US (almost 400 cases) per law firm advertisement.
http://www.youtube.com/watch?NR=1&v=-Y74iV5dUZQ
Louisiana Tech University
College of Engineering and Science

## Slide 26: Stryker hip implant
Figure http://www.millerandzois.com/stryker-hip-settlement.html
“My dad’s hip is squeaking!” http://www.youtube.com/watch?NR=1&v=dkLDm9tgW1U
Man walking up stairs “T This is the sound a $40,000.00 Stryker Ceramic on Ceramic hip makes.”
http://www.youtube.com/watch?v=rviIIuGC1Xw
Louisiana Tech University
College of Engineering and Science

## Slide 27: Timeline of Stryker recall
February 3, 2003: FDA approves hip replacement system.
January 2005 to April 2007: Stryker receives large number of patient complaints, but takes no significant action.
March 15, 2007: FDA sends warning letter about manufacturing facility in Ireland. “… inspection revealed that these devices are adulterated within the meaning of section 501(h) of the Act (21 U.S.C. § 351(h)), in that the methods used in, or the facilities or controls used for, their manufacture, packing, storage, or installation are not in conformity with the Current Good Manufacturing Practice (CGMP) requirements of the Quality System (QS) regulation.”
Source: http://www.millerandzois.com/stryker-hip-settlement.html
Louisiana Tech University
College of Engineering and Science

## Slide 28: Timeline of Stryker recall
November 28, 2007: FDA sends second letter warning Stryker. “Customer complaints have been received for squeaking/noisy ceramic on ceramic hip joints.”
January 22, 2008: Stryker initiates a voluntary recall. “As a result of the deviation from internal specifications, the Company is initiating a voluntary recall of Trident PSL and Hemispherical Acetabular Cups manufactured in its Cork facility.”
Source: http://www.millerandzois.com/stryker-hip-settlement.html
Source: http://www.millerandzois.com/stryker-hip-settlement.html
Louisiana Tech University
College of Engineering and Science

## Slide 29: Hip Implant recalls
Depuy ASR XL Acetabular and ASR Hip Resurfacing Systems (metal-on-metal head and cup)
Smaller cup than other mfgrs.  edge-loading, metal debris, loosened implants (osteolysis?)
Zimmer Durom Cup (cup has porous surface for bone ingrowth)
No bone ingrowth; loosened; severe pain
Stryker Trident PSL and Trident Hemispherical Acetubular Cup (ceramic cups)
Squeaks and pops
Louisiana Tech University
College of Engineering and Science

## Slide 30: Payments to Doctors
A study looked at payments to doctors from 5 orthopedic device companies (Archives of Internal Medicine, Sept 2010). In 2007, 41 doctors were paid $114 million, over 1 million each. Of these, 32 published papers the next year and half did not mention the large payments.
The companies were: *DePuy, *Stryker, *Zimmer, Biomet, and Smith & Nephew (*these also had product recalls, some noted previously)
Is doctor reports on a company’s products in a scholarly, scientific journal that fails to mention being paid $1 million dollars in a single year by that same company ethical?
Louisiana Tech University
College of Engineering and Science
2010 http://www.hip-replacement-and-recovery.com/payments-to-doctors.html

## Slide 31: Optional: Types of hip implants
You do not need to know this for an exam!

## Slide 32: Main types of hip implants
Louisiana Tech University
College of Engineering and Science
Total hip arthroplasty
  Cemented cup and stem
  Cementless (press fit) cup and stem
  Hybrid (cemented stem, press fit cup )
Partial replacement
  Femoral head only
  Acetabular cup only
Hip resurfacing
  Cover (resurface) femoral head
  Replacement cup

## Slide 33
http://emedicine.medscape.com/article/398669-overview
Total hip arthroplasty
Femoral component (metal, press fit, cementless) and acetabular component (polyethylene cup with metal backing, screw fixation)
Louisiana Tech University
College of Engineering and Science

## Slide 34
http://emedicine.medscape.com/article/398669-overview
Total hip arthroplasty
Femoral component (cemented metal)
Acetabular cup (polyethylene, cemented [yellow arrow points to radio-opaque rim of cup])
Louisiana Tech University
College of Engineering and Science

## Slide 35
http://emedicine.medscape.com/article/398669-overview
Total hip arthroplasty
Radiograph shows femoral component (ceramic head [arrow], metal stem [no cement], cable) and acetabular component (ceramic cup, metal backing, cementless)
Note: ceramic head is less dense than metal
Louisiana Tech University
College of Engineering and Science

## Slide 36
Resurfacing total hip arthroplasty.
Femoral component (metal head and stem) Acetabulum component (metal, cementless).
Note: large femoral head and preserved femoral neck.
http://emedicine.medscape.com/article/398669-overview
Louisiana Tech University
College of Engineering and Science

## Slide 37
Regenerex porous titanium fixation for acetabular cup
http://www.youtube.com/watch?v=d18mN_3SAZI&feature=related
Louisiana Tech University
College of Engineering and Science